Outstanding problems.

- Does CN pump need rebuilt?

- What was the battery report? Who is maintaining the battery log.

- Dilution valve still needs attention. Here is dilution valve operation recent history

  - 898 – worked shifted at 217 hPa (1 km free fall)

  - 899 – not plugged in, so CN blocked

  - 900 – not flown, no CN

  - 901 – not flown, dual CN

  - 902 – did not shift out of dilution

  - 903 – not flown, no CN

  - 904 – worked shifted at 213 hPa

  - 905 – not flown, dual CN

  - 906 – was not in dilution ever

  - 907 – was not in dilution ever

  - 908 – worked, shifted at 208 (236) hPa ascent (descent). Dilution flow = 0.5 cc/s.

- Spares need to be maintained at the launch shack of: microprocessor, power supplies, Vaisala PTU sensor, 400 MHz Xmitter – done?

- Jupiter dropouts for ls and rv files – better but still noisey.

- Paper over the non-working lift scale, or find out why it doesn’t work.

Flight by Flight

Wy908 - 130326

Preflight

- 400 MHz receiver failed – power supply?

- 400 MHz preamp hanging by one wire. Preamps not labeled in and out.

- Batteries not charged and ready for replacement.

- FindU balloon call sign had been removed from findU web site.

Day of flight

- No internet on one computer – fix = turn off modem for 15 sec

- Anemometer not reliable

- Balloon twisted as it was inflated.

- Inflation tube right near the limit for a 30 ft bubble.

- Final strap tightening and buttoning up on gondola was not completed.

- Yellow wrapped tightly to balloon in middle section – in reducing this compression a hole was torn in the balloon – fixed with balloon tape.

- Surface winds too strong at release. Sunrise lull did not last.

- Spot does not work above 60,000 ft.

- New spot (stratuw) does not report as frequently as ASCII.

- New GPS mapping software locking up new recovery computer.

Kudos

- Found the balloon tape.

- Ran hard and no pendulum into ground

- Spot gave good final points on balloon and gondola.

- Were patient to wait for a lull.

Wy907 - 130117

Preflight

- One trakpak computer froze long delay

- Yellow wrap very tight to balloon and base of balloon vacuum packed?

- Instruments out to early

In flight

- Balloon failure 121 hPa.

- Dilution valve did not work

- Batteries CN – flashing, Jup 1/5, Transp 4/5

Kudos

- 30 foot bubble good

- Gondola suspension worked well

- Uplooking J3m FW antenna better, but still Jupiter ls and rv files noisey.

- Great to have uplooking video of balloon.

- Hand held pivot at launch to handle wind shift worked well

Wy906

Preflight

- Incorrect balloon weight

- No anemometer

- Primary command only triggers when secondary is on.

- Spot not outside when turned on

- 3.5 tubes he left

- Dropper upside down

- 400 MHz preamp not plugged in

  - Antenna cable came apart

  - FM band got switched to FMn

In flight

- Line caught on j3m sunshield

- J3m tm not good

- Balloon failure at 22 hPa

- Dilution valve did not work

Wy905-120920, j3m, ms38n, cn40, cn34h

Preflight:

- Forgot flight book and data medium

- Forgot to check previous flight problems.

- Internet connection was lost on gs computer – bad internet card.

- CN µprocessor bad and took a long time to isolate.

- Primary command radio failed after successful test in the lab.

- Did not complete the preflight run through once we found problems with the internet connection.

- White straps from pervious flights not stored where they could be reused.

- Spot not on check list.

In flight:

- 2nd 400 MHz preamp not plugged in.

- 6 hr trajectory significantly east of impact point – probably due to light pressure gradient at upper levels.

- Uplooking camera didn’t start

- Jupiter ozone reporting numbers > 1262

- FW telemetry for J3m data is still not acceptable.

wy901-120322, j3m, ms38v, cn40, cn34h

Preflight:

- Forgot command system flight bag

- Forgot strap for chute cut-off

- Transponder failed to report altitude – lead to cancelled flight in early March.

- Vaisala numbers not recorded on instrument set-up sheets

- CN flow rates not on set up sheets

- Jupiter flow rate not included on instrument

- Interior blocks prepped ahead of time

- CN heater and stack straps not prepared ahead

- Jupiter higher channels not reporting initially.

- A little late in getting inflation started.

- In calm wind situation, should be more patient and hold piball if there is a wind shift to insure it is permanent.

- Could prepare inflation tube when loading balloon when they are so long.

Flight:

- Jupiter higher channels intermittant reporting, data okay.

- Jupiter fw drop-outs still not understood.

- Preamp on 2nd 403 antenna worked well.

- Heater onCN34h did not increase at lower pressures.

- Jupiter ozone not good, too low. (where was the 18.5 V battery?)

- Tangled parachute

- Jupiter pump got quite warm.

Recovery/post flight:

- Recovery and loadline equipment boxes not returned to main building.

- Failed to have right media for data transfer from the systems at launch shack.

Post-flight fixes:

- Shroud lines on 28 ft chute needs some work.

- Jupiter fw

- Heater on cn34

- Jupiter pump needs rebuilt