## Cal_points_May 2011
Last modified May 2011 (td) |  |  |  | Calibration constants from new poly in Blue | [6 empty] | Conversion | Used
for NEW Stan's new PHA | [6 empty] | New PHA SS |  |  |  | on Compaq | briefly
Date = | ,110510 |  |  |  | slope | m | 128.3 | 128.87 | 128 |  | 126.8 | 133
[5 empty] | offset | b | -11.51 | -14.67 | 0 |  | -46.7 | -43.6
Counter | Rmax on | CR Mie | Poly | Poly | CR Mie | Voltage | Channel numbers |  |  | CR Calc | Channel Numbers |  |  | Channel | CntrRespon | Disc Level | CntrRespon | Disc Level
Board | board (um) | m=1.45 | Diam | Radius | m=1.59 |  | mV+b | mV+b | mV+b | Source | mV+b | mV+b | XaCh | Radius | m=1.45 | V | m=1.45 | V
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_Vol | 12 Chan | 40 Degrees | New Bulb May 2006 |  |  |  | w/o 1V pnt | w 1V pnt | Linear |  | QunMiao2 | QunMiao0 |  |  | File-nuwy_opc.dat |  | File=40253159.dat |  | Act DL | Implied CR
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 241.4976 | 239.4616 | 252.416 | nuwy_opc.dat | 203.3496 | 218.676 | 1 | 0.15 | 4.94 | 0.1976 | 4.92 | 0.18743 | 0.08 | 2
| 1.08 | 200 | 0.89 | 0.445 | 48.2 | 1.928 | 235.8524 | 233.7914 | 246.784 | nuwy_opc.dat | 197.7704 | 212.824 | 3 | 0.25 | 30.4 | 1.216 | 30.3 | 1.1543 | 1.1 | 27.5
| 1.08 | 210.4 | 0.89 | 0.445 | 47.4 | 1.8023 | 219.7227 | 217.59 | 230.692 | 40253159.dat | 181.8293 | 196.1034 | 6 | 0.49 | 72.3 | 2.892 | 71.9 | 2.739 | 2.73 | 68.25
[13 empty] | 9 | 1.08 | 200 | 8 | 210 | 8 | 7.999 | 199.975
[12 empty] | Dynamic Range |  |  | 40.4858
Mid Gain | 1.25 | 304 | 1.11 | 0.555 | 59.75 | 1.5724 | 190.2249 | 187.9611 | 201.2632 | nuwy_opc.dat | 152.6763 | 165.525 | 2 | 0.19 | 13.3 | 0.35 | 13.3 | 0.33885 | 0.35 | 13.3
| 1.25 | 304 | 2.02 | 1.01 | 214 | 5.6316 | 711.0216 | 711.0716 | 720.8421 | nuwy_opc.dat | 667.3842 | 705.4 | 4 | 0.3 | 39.7 | 1.0447 | 39.5 | 1.0064 | 1.04 | 39.52
| 1.25 | 314 | 1.11 | 0.555 | 57.95 | 1.4764 | 177.9164 | 175.5979 | 188.9834 | 40253159.dat | 140.5117 | 152.7656 | 7 | 0.62 | 103 | 2.7105 | 103 | 2.6242 | 2.71 | 102.98
[13 empty] | 10 | 1.25 | 304 | 8 | 314 | 8 | 8 | 304
[12 empty] | Dynamic Range |  |  | 22.8571
Low Gain | 2 | 702 | 1.992 | 0.996 | 210.4 | 2.3977 | 296.1176 | 294.3243 | 306.9083 | nuwy_opc.dat | 257.331 | 275.2969 | 5 | 0.38 | 46 | 0.52422 | 46 | 0.51685 | 0.52 | 45.63
| 2 | 702 | 2.02 | 1.01 | 214 | 2.4387 | 301.3812 | 299.6113 | 312.1595 | nuwy_opc.dat | 262.533 | 280.7533 | 8 | 0.78 | 114 | 1.2991 | 114 | 1.2809 | 1.3 | 114.075
| 2 | 711.6 | 2.02 | 1.01 | 208 | 2.3384 | 288.5057 | 286.6786 | 299.3142 | 40253159.da | 249.8082 | 267.4062 | 11 | 1.58 | 487 | 5.5499 | 497 | 5.5843 | 5.55 | 487.0125
[13 empty] | 12 | 2 | 702 | 8 | 712 | 8 | 8 | 702
[12 empty] | Dynamic Range |  |  | 15.2609
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_Vol | 12 Chan | 40 Degrees | [12 empty] | File-nuwy_opc.dat |  | File=40253159.dat |  | Act DL | Implied CR
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 241.4976 | 239.4616 | 252.416 | nuwy_opc.dat | 203.3496 | 218.676 |  | 0.15 | 4.94 | 0.1976 | 4.92 | 0.18743 | 0.08 | 2
| 1.08 | 200 | 0.89 | 0.445 | 48.2 | 1.928 | 235.8524 | 233.7914 | 246.784 | nuwy_opc.dat | 197.7704 | 212.824 |  | 0.25 | 30.4 | 1.216 | 30.3 | 1.1543 | 1.1 | 27.5
| 1.08 | 210.4 | 0.89 | 0.445 | 47.4 | 1.8023 | 219.7227 | 217.59 | 230.692 | 40253159.dat | 181.8293 | 196.1034 |  | 0.49 | 72.3 | 2.892 | 71.9 | 2.739 | 2.73 | 68.25
[14 empty] | 1.08 | 200 | 8 | 210 | 8 | 7.999 | 199.975
[12 empty] | Dynamic Range |  |  | 40.4858
Mid Gain | 0.62 | 103 | 0.701 | 0.3505 | 38.1 | 2.9592 | 368.1583 | 366.6851 | 378.7806 | nuwy_opc.dat | 328.5295 | 349.9767 |  | 0.19 | 13.3 | 1.033 | 13.3 | 1.033 | 1.03 | 13.2613
| 0.62 | 103 | 0.653 | 0.3265 | 39 | 3.0291 | 377.1269 | 375.6935 | 387.7282 | nuwy_opc.dat | 337.3932 | 359.2738 |  | 0.3 | 39.7 | 3.0835 | 39.5 | 3.068 | 3.06 | 39.3975
| 0.62 | 103.4 | 0.653 | 0.3265 | 38.8 | 3.0019 | 373.6382 | 372.1893 | 384.2476 | 40253159.dat | 333.9453 | 355.6573 |  | 0.38 | 46 | 3.5728 | 46 | 3.5728 | 3.56 | 45.835
[14 empty] | 0.62 | 103 | 8 | 103 | 8 | 8 | 103
[12 empty] | Dynamic Range |  |  | 7.7444
| 0.62 | 103.4 | 0.624 | 0.312 | 39.3 | 3.0406 | 378.6014 | 377.1746 | 389.1992 |  | 338.8505 | 360.8023
Low Gain | 2 | 702 | 1.992 | 0.996 | 210.4 | 2.3977 | 296.1176 | 294.3243 | 306.9083 | nuwy_opc.dat | 257.331 | 275.2969 |  | 0.78 | 114 | 1.2991 | 114 | 1.2809 | 1.28 | 112.32
| 2 | 702 | 2.02 | 1.01 | 214 | 2.4387 | 301.3812 | 299.6113 | 312.1595 | nuwy_opc.dat | 262.533 | 280.7533 |  | 1.25 | 304 | 3.4644 | 314 | 3.5281 | 3.53 | 309.7575
| 2 | 711.6 | 2.02 | 1.01 | 208 | 2.3384 | 288.5057 | 286.6786 | 299.3142 | 40253159.da | 249.8082 | 267.4062 |  | 1.58 | 487 | 5.5499 | 497 | 5.5843 | 5.59 | 490.5225
[14 empty] | 2 | 702 | 8 | 712 | 8 | 8 | 702
[12 empty] | Dynamic Range |  |  | 6.1579
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_Vol | 2 Chan | 27 Degrees | [12 empty] | File=27953159.dat |  | Revised, 110503 | File=27953159.dat
High Gain | 0.65 | 104 | 0.903 | 0.4515 | 52.53 | 4.0408 | 506.9207 | 506.0639 | 517.2185 | 27953159.dat | 465.6695 | 493.8223 |  | 0.15 | 5.31 | 0.4818 | 0.15 | 8.83 | 0.67701 | New analysis
| 0.25 | 51.8 | 0.89 | 0.445 | 53.9 | 8.3243 | 1056.5008 | 1058.0857 | 1065.5135 | 27953159.dat | 1008.8243 | 1063.5351 |  | 0.25 | 51.8 | 4.7 | 0.25 | 61.3 | 4.7 | set points
[15 empty] | 88.1702 | 8 | 0.65 | 104.3404 | 8 | May, 2011
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_PSC | 12 Chan | 40 Degrees | [12 empty] | File-nuwy_opc.dat |  | File=40253159.dat
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 241.4976 | 239.4616 | 252.416 | nuwy_opc.dat | 203.3496 | 218.676 |  | 0.15 | 4.94 | 0.1976 | 4.92 | 0.18743
| 1.08 | 200 | 0.89 | 0.445 | 48.2 | 1.928 | 235.8524 | 233.7914 | 246.784 | nuwy_opc.dat | 197.7704 | 212.824 |  | 0.25 | 30.4 | 1.216 | 30.3 | 1.1543
| 1.08 | 210.4 | 0.89 | 0.445 | 47.4 | 1.8023 | 219.7227 | 217.59 | 230.692 | 40253159.dat | 181.8293 | 196.1034 |  | 0.49 | 72.3 | 2.892 | 71.9 | 2.739
[14 empty] | 1.08 | 200 | 8 | 210 | 8
[12 empty] | Dynamic Range |  |  | 40.4858
Mid Gain | 1.75 | 565 | 1.992 | 0.996 | 210.4 | 2.9791 | 370.7105 | 369.2486 | 381.3267 | nuwy_opc.dat | 331.0518 | 352.6223 |  | 0.3 | 39.7 | 0.56212 | 39.5 | 0.55052
| 1.75 | 565 | 2.02 | 1.01 | 214 | 3.0301 | 377.2504 | 375.8175 | 387.8513 | nuwy_opc.dat | 337.5152 | 359.4018 |  | 0.75 | 109 | 1.5434 | 110 | 1.5331
| 1.75 | 574 | 2.02 | 1.01 | 208 | 2.899 | 360.4259 | 358.9183 | 371.0662 | 40253159.dat | 320.8875 | 341.961 |  | 1.25 | 304 | 4.3044 | 314 | 4.3763
[14 empty] | 1.75 | 565 | 8 | 574 | 8
[12 empty] | Dynamic Range |  |  | 14.2317
Low Gain | 10 | 16900 | 6.992 | 3.496 | 2195 | 1.0391 | 121.8005 | 119.2328 | 132.9988 | nuwy_opc.dat | 85.052 | 94.5941 |  | 2.5 | 1140 | 0.53964 | 1140 | 0.54286
| 10 | 16900 | 7 | 3.5 | 2210 | 1.0462 | 122.7115 | 120.1478 | 133.9077 | nuwy_opc.dat | 85.9523 | 95.5385 |  | 3.5 | 2170 | 1.0272 | 2170 | 1.0333
| 10 | 16880 | 7 | 3.5 | 2120 | 1.0047 | 117.3981 | 114.8108 | 128.6066 | 40253159.dat | 80.7009 | 90.0303 |  | 5 | 4300 | 2.0355 | 4310 | 2.0524
[14 empty] | 10 | 16900 | 8 | 16800 | 8
[12 empty] | Dynamic Range |  |  | 14.8246
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
[9 empty] | 1024 |  | 1024 | 1024 | 512
CN | 4 Chan | 25 Degrees | [6 empty] | Scale |  | Scale | Scale | Scale |  | File=ae253159.dat
| 0.8 | 186 | 0.903 | 0.4515 | 68.1 | 2.929 | 364.2848 | 362.7944 | 374.9161 |  | 324.7013 | 345.9613 | 172.9806 | 0.15 | 6.04 | 0.25978
| 0.8 | 186 | 0.89 | 0.445 | 70 | 3.0108 | 374.7696 | 373.3257 | 385.3763 |  | 335.0634 | 356.8301 | 178.4151 | 0.25 | 59.8 | 2.572
| 0.8 | 186 | 1.011 | 0.5055 | 65 | 2.7957 | 347.1782 | 345.6117 | 357.8495 |  | 307.7946 | 328.228 | 164.114 | 0.3 | 105 | 4.5161
[14 empty] | 0.8 | 186 | 8
| 0.8 | 186 | 0.945 | 0.4725 | 64 | 2.7527 | 341.6599 | 340.0689 | 352.3441 |  | 302.3409 | 322.5075
CN | 3 Chan | 25 Degrees | [12 empty] | File=ae253159.dat
| 0.3 | 105 | 0.903 | 0.4515 | 68.1 | 5.1886 | 654.1837 | 653.9812 | 664.1371 |  | 611.2109 | 646.48 | 323.24 | 0.15 | 6.04 | 0.46019
| 0.3 | 105 | 0.89 | 0.445 | 70 | 5.3333 | 672.7567 | 672.6367 | 682.6667 |  | 629.5667 | 665.7333 | 332.8667 | 0.25 | 59.8 | 4.5562
| 0.3 | 105 | 1.011 | 0.5055 | 65 | 4.9524 | 623.8805 | 623.5433 | 633.9048 |  | 581.2619 | 615.0667 | 307.5333 | 0.3 | 105 | 8
| 0.3 | 105 | 0.945 | 0.4725 | 64 | 4.8762 | 614.1052 | 613.7247 | 624.1524 |  | 571.601 | 604.9333 | 302.4667
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
[9 empty] | 1024 |  | 1024 | 1024 | 512
Dust | 4 Chan | 25 Degrees | ,----------------------------------------------- | [5 empty] | Scale |  | Scale | Scale | Scale |  | File=ae253159.dat
| 0.3 | 105 | 0.903 | 0.4515 | 68.1 | 5.1886 | 654.1837 | 653.9812 | 664.1371 |  | 611.2109 | 646.48 | 323.24 | 0.15 | 6.04 | 0.46019
| 0.3 | 105 | 0.89 | 0.445 | 70 | 5.3333 | 672.7567 | 672.6367 | 682.6667 |  | 629.5667 | 665.7333 | 332.8667 | 0.2 | 23.7 | 1.8057
| 0.3 | 105 | 1.011 | 0.5055 | 65 | 4.9524 | 623.8805 | 623.5433 | 633.9048 |  | 581.2619 | 615.0667 | 307.5333 | 0.25 | 59.8 | 4.5562
| 0.3 | 105 | 0.945 | 0.4725 | 64 | 4.8762 | 614.1052 | 613.7247 | 624.1524 |  | 571.601 | 604.9333 | 302.4667 | 0.3 | 105 | 8
[6 empty] | 8 | 1014.89
Counter response values checked July 5, 2011, using Compare_Mie_OPC.pro, rfr,opcangle and print_cr for Dust and CN

## Poly_Sizes
PolySizes | Available(Dia) | 0.261 | 0.364 | 0.5 | 0.653 | 0.89 | 0.913 | 0.721 | 0.804 | 0.945 | 1.011 | 1.09 | 1.11 | 2.02 | 2.92 | 3.4 | 4 | 4.92 | 7.04
| Nebulized |  | Y | Y | Y | Y |  |  |  | Y | Y |  | Y | Y | Y | Y | Y | Y | Y
Poly Sizes | New | April, 2003 |  |  | 0.7 | 0.903 | 1.992 | 6.992

## Cal_Points_020218
Last modified May 2011 (td) |  |  |  | Calibration constants from new poly in Blue | [5 empty] | Calibration constants to use in bold
Counter | Max Radii | Cntr Respons | Poly | Poly | CntrRespon | Voltage | Channel# |  | Channel# |  | CR Calc |  | Channel#
Board | Meas by | Mie Calc | Diam | Radius | Mie Calc |  | QunMiao2 = |  | New PHA_SS |  | Source |  | QunMiao =
| Board (um) | m-1.45 |  |  | m=1.59 |  | 126.8*V-46.7 |  | 128*V |  |  |  | 133*V-43.6
[7 empty] | Conversion |  | New Current |  |  |  | Used only |  | Channel | CntrRespon | Disc Level | CntrRespon | Disc Level
[7 empty] | used on Compaq |  | Convesion |  |  |  | briefly | XaCh | Radius | m=1.45 | V | m=1.45 | V
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_Vol | 12 Chan | 40 Degrees | New Bulb May 2006 | [12 empty] | File-nuwy_opc.dat |  | File=40253159.dat |  | Act DL | Implied CR
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 203.3496 |  | 252.416 |  | nuwy_opc.dat |  | 218.676 | 1 | 0.15 | 4.94 | 0.1976 | 4.92 | 0.18743 | 0.08 | 2
| 1.08 | 200 | 0.89 | 0.445 | 48.2 | 1.928 | 197.7704 |  | 246.784 |  | nuwy_opc.dat |  | 212.824 | 3 | 0.25 | 30.4 | 1.216 | 30.3 | 1.1543 | 1.1 | 27.5
| 1.08 | 210.4 | 0.89 | 0.445 | 47.4 | 1.8023 | 181.8293 |  | 230.692 |  | 40253159.dat |  | 196.1034 | 6 | 0.49 | 72.3 | 2.892 | 71.9 | 2.739 | 2.73 | 68.25
[14 empty] | 9 | 1.08 | 200 | 8 | 210 | 8 | 7.999 | 199.975
[13 empty] | Dynamic Range |  |  | 40.4858
Mid Gain | 1.25 | 304 | 1.11 | 0.555 | 59.75 | 1.5724 | 152.6763 |  | 201.2632 |  | nuwy_opc.dat |  | 165.525 | 2 | 0.19 | 13.3 | 0.35 | 13.3 | 0.33885 | 0.35 | 13.3
| 1.25 | 304 | 2.02 | 1.01 | 214 | 5.6316 | 667.3842 |  | 720.8421 |  | nuwy_opc.dat |  | 705.4 | 4 | 0.3 | 39.7 | 1.0447 | 39.5 | 1.0064 | 1.04 | 39.52
| 1.25 | 314 | 1.11 | 0.555 | 57.95 | 1.4764 | 140.5117 |  | 188.9834 |  | 40253159.dat |  | 152.7656 | 7 | 0.62 | 103 | 2.7105 | 103 | 2.6242 | 2.71 | 102.98
[14 empty] | 10 | 1.25 | 304 | 8 | 314 | 8 | 8 | 304
[13 empty] | Dynamic Range |  |  | 22.8571
Low Gain | 2 | 702 | 1.992 | 0.996 | 210.4 | 2.3977 | 257.331 |  | 306.9083 |  | nuwy_opc.dat |  | 275.2969 | 5 | 0.38 | 46 | 0.52422 | 46 | 0.51685 | 0.52 | 45.63
| 2 | 702 | 2.02 | 1.01 | 214 | 2.4387 | 262.533 |  | 312.1595 |  | nuwy_opc.dat |  | 280.7533 | 8 | 0.78 | 114 | 1.2991 | 114 | 1.2809 | 1.3 | 114.075
| 2 | 711.6 | 2.02 | 1.01 | 208 | 2.3384 | 249.8082 |  | 299.3142 |  | 40253159.da |  | 267.4062 | 11 | 1.58 | 487 | 5.5499 | 497 | 5.5843 | 5.55 | 487.0125
[14 empty] | 12 | 2 | 702 | 8 | 712 | 8 | 8 | 702
[13 empty] | Dynamic Range |  |  | 15.2609
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_Vol | 12 Chan | 40 Degrees | [13 empty] | File-nuwy_opc.dat |  | File=40253159.dat |  | Act DL | Implied CR
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 203.3496 |  | 252.416 |  | nuwy_opc.dat |  | 218.676 |  | 0.15 | 4.94 | 0.1976 | 4.92 | 0.18743 | 0.08 | 2
| 1.08 | 200 | 0.89 | 0.445 | 48.2 | 1.928 | 197.7704 |  | 246.784 |  | nuwy_opc.dat |  | 212.824 |  | 0.25 | 30.4 | 1.216 | 30.3 | 1.1543 | 1.1 | 27.5
| 1.08 | 210.4 | 0.89 | 0.445 | 47.4 | 1.8023 | 181.8293 |  | 230.692 |  | 40253159.dat |  | 196.1034 |  | 0.49 | 72.3 | 2.892 | 71.9 | 2.739 | 2.73 | 68.25
[15 empty] | 1.08 | 200 | 8 | 210 | 8 | 7.999 | 199.975
[13 empty] | Dynamic Range |  |  | 40.4858
Mid Gain | 0.62 | 103 | 0.701 | 0.3505 | 38.1 | 2.9592 | 328.5295 |  | 378.7806 |  | nuwy_opc.dat |  | 349.9767 |  | 0.19 | 13.3 | 1.033 | 13.3 | 1.033 | 1.03 | 13.2613
| 0.62 | 103 | 0.653 | 0.3265 | 39 | 3.0291 | 337.3932 |  | 387.7282 |  | nuwy_opc.dat |  | 359.2738 |  | 0.3 | 39.7 | 3.0835 | 39.5 | 3.068 | 3.06 | 39.3975
| 0.62 | 103.4 | 0.653 | 0.3265 | 38.8 | 3.0019 | 333.9453 |  | 384.2476 |  | 40253159.dat |  | 355.6573 |  | 0.38 | 46 | 3.5728 | 46 | 3.5728 | 3.56 | 45.835
[15 empty] | 0.62 | 103 | 8 | 103 | 8 | 8 | 103
[13 empty] | Dynamic Range |  |  | 7.7444
| 0.62 | 103.4 | 0.624 | 0.312 | 39.3 | 3.0406 | 338.8505 |  | 389.1992 |  |  |  | 360.8023
Low Gain | 2 | 702 | 1.992 | 0.996 | 210.4 | 2.3977 | 257.331 |  | 306.9083 |  | nuwy_opc.dat |  | 275.2969 |  | 0.78 | 114 | 1.2991 | 114 | 1.2809 | 1.28 | 112.32
| 2 | 702 | 2.02 | 1.01 | 214 | 2.4387 | 262.533 |  | 312.1595 |  | nuwy_opc.dat |  | 280.7533 |  | 1.25 | 304 | 3.4644 | 314 | 3.5281 | 3.53 | 309.7575
| 2 | 711.6 | 2.02 | 1.01 | 208 | 2.3384 | 249.8082 |  | 299.3142 |  | 40253159.da |  | 267.4062 |  | 1.58 | 487 | 5.5499 | 497 | 5.5843 | 5.59 | 490.5225
[15 empty] | 2 | 702 | 8 | 712 | 8 | 8 | 702
[13 empty] | Dynamic Range |  |  | 6.1579
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_Vol | 2 Chan | 27 Degrees | [13 empty] | File=27953159.dat |  | Revised, 110503 | File=27953159.dat
High Gain | 0.65 | 104 | 0.903 | 0.4515 | 52.53 | 4.0408 | 465.6695 |  | 517.2185 |  | 27953159.dat |  | 493.8223 |  | 0.15 | 5.31 | 0.4818 | 0.15 | 8.83 | 0.67701 | New analysis of the set points in May 2011
| 0.25 | 51.8 | 0.89 | 0.445 | 53.9 | 4.8905 | 573.4205 |  | 625.9892 |  | 27953159.dat |  | 606.8419 |  | 0.25 | 51.8 | 4.7 | 0.25 | 61.3 | 4.7
[16 empty] | 88.1702 | 8 | 0.65 | 104.3404 | 8
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_PSC | 12 Chan | 40 Degrees | [13 empty] | File-nuwy_opc.dat |  | File=40253159.dat
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 203.3496 |  | 252.416 |  | nuwy_opc.dat |  | 218.676 |  | 0.15 | 4.94 | 0.1976 | 4.92 | 0.18743
| 1.08 | 200 | 0.89 | 0.445 | 48.2 | 1.928 | 197.7704 |  | 246.784 |  | nuwy_opc.dat |  | 212.824 |  | 0.25 | 30.4 | 1.216 | 30.3 | 1.1543
| 1.08 | 210.4 | 0.89 | 0.445 | 47.4 | 1.8023 | 181.8293 |  | 230.692 |  | 40253159.dat |  | 196.1034 |  | 0.49 | 72.3 | 2.892 | 71.9 | 2.739
[15 empty] | 1.08 | 200 | 8 | 210 | 8
[13 empty] | Dynamic Range |  |  | 40.4858
Mid Gain | 1.75 | 565 | 1.992 | 0.996 | 210.4 | 2.9791 | 331.0518 |  | 381.3267 |  | nuwy_opc.dat |  | 352.6223 |  | 0.3 | 39.7 | 0.56212 | 39.5 | 0.55052
| 1.75 | 565 | 2.02 | 1.01 | 214 | 3.0301 | 337.5152 |  | 387.8513 |  | nuwy_opc.dat |  | 359.4018 |  | 0.75 | 109 | 1.5434 | 110 | 1.5331
| 1.75 | 574 | 2.02 | 1.01 | 208 | 2.899 | 320.8875 |  | 371.0662 |  | 40253159.dat |  | 341.961 |  | 1.25 | 304 | 4.3044 | 314 | 4.3763
[15 empty] | 1.75 | 565 | 8 | 574 | 8
[13 empty] | Dynamic Range |  |  | 14.2317
Low Gain | 10 | 16900 | 6.992 | 3.496 | 2195 | 1.0391 | 85.052 |  | 132.9988 |  | nuwy_opc.dat |  | 94.5941 |  | 2.5 | 1140 | 0.53964 | 1140 | 0.54286
| 10 | 16900 | 7 | 3.5 | 2210 | 1.0462 | 85.9523 |  | 133.9077 |  | nuwy_opc.dat |  | 95.5385 |  | 3.5 | 2170 | 1.0272 | 2170 | 1.0333
| 10 | 16880 | 7 | 3.5 | 2120 | 1.0047 | 80.7009 |  | 128.6066 |  | 40253159.dat |  | 90.0303 |  | 5 | 4300 | 2.0355 | 4310 | 2.0524
[15 empty] | 10 | 16900 | 8 | 16800 | 8
[13 empty] | Dynamic Range |  |  | 14.8246
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
[7 empty] | 1024 | 512 | 1024 | 512 |  |  | 1024 | 512
CN | 4 Chan | 25 Degrees |  |  |  |  | Scale | Scale | Scale | Scale |  |  | Scale | Scale |  | File=ae253159.dat
| 0.8 | 186 | 0.903 | 0.4515 | 68.1 | 2.929 | 324.7013 | 162.3506 | 374.9161 | 187.4581 |  |  | 345.9613 | 172.9806 | 0.15 | 6.04 | 0.25978
| 0.8 | 186 | 0.89 | 0.445 | 70 | 3.0108 | 335.0634 | 167.5317 | 385.3763 | 192.6882 |  |  | 356.8301 | 178.4151 | 0.25 | 59.8 | 2.572
| 0.8 | 186 | 1.011 | 0.5055 | 65 | 2.7957 | 307.7946 | 153.8973 | 357.8495 | 178.9247 |  |  | 328.228 | 164.114 | 0.3 | 105 | 4.5161
[15 empty] | 0.8 | 186 | 8
| 0.8 | 186 | 0.945 | 0.4725 | 64 | 2.7527 | 302.3409 | 151.1704 | 352.3441 | 176.172 |  |  | 322.5075 | 161.2538
CN | 3 Chan | 25 Degrees | [13 empty] | File=ae253159.dat
| 0.3 | 105 | 0.903 | 0.4515 | 68.1 | 5.1886 | 611.2109 | 305.6054 | 664.1371 | 332.0686 |  |  | 646.48 | 323.24 | 0.15 | 6.04 | 0.46019
| 0.3 | 105 | 0.89 | 0.445 | 70 | 5.3333 | 629.5667 | 314.7833 | 682.6667 | 341.3333 |  |  | 665.7333 | 332.8667 | 0.25 | 59.8 | 4.5562
| 0.3 | 105 | 1.011 | 0.5055 | 65 | 4.9524 | 581.2619 | 290.631 | 633.9048 | 316.9524 |  |  | 615.0667 | 307.5333 | 0.3 | 105 | 8
| 0.3 | 105 | 0.945 | 0.4725 | 64 | 4.8762 | 571.601 | 285.8005 | 624.1524 | 312.0762 |  |  | 604.9333 | 302.4667
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
[7 empty] | 1024 | 512 | 1024 | 512 |  |  | 1024 | 512
Dust | 4 Chan | 25 Degrees | ,----------------------------------------------- |  |  |  | Scale | Scale | Scale | Scale |  |  | Scale | Scale |  | File=ae253159.dat
| 0.3 | 105 | 0.903 | 0.4515 | 68.1 | 5.1886 | 611.2109 | 305.6054 | 664.1371 | 332.0686 |  |  | 646.48 | 323.24 | 0.15 | 6.04 | 0.46019
| 0.3 | 105 | 0.89 | 0.445 | 70 | 5.3333 | 629.5667 | 314.7833 | 682.6667 | 341.3333 |  |  | 665.7333 | 332.8667 | 0.2 | 23.7 | 1.8057
| 0.3 | 105 | 1.011 | 0.5055 | 65 | 4.9524 | 581.2619 | 290.631 | 633.9048 | 316.9524 |  |  | 615.0667 | 307.5333 | 0.25 | 59.8 | 4.5562
| 0.3 | 105 | 0.945 | 0.4725 | 64 | 4.8762 | 571.601 | 285.8005 | 624.1524 | 312.0762 |  |  | 604.9333 | 302.4667 | 0.3 | 105 | 8
PolySizes | Available(Dia) | 0.261 | 0.364 | 0.5 | 0.653 | 0.89 | 0.913 | 0.721 | 0.804 | 0.945 | 1.011 | 1.09 | 1.11 | 2.02 | 2.92 | 3.4 | 4 | 4.92 | 7.04
| Nebulized |  | Y | Y | Y | Y |  |  |  | Y | Y |  | Y | Y | Y | Y | Y | Y | Y
Poly Sizes | New | April, 2003 |  |  | 0.7 |  |  |  |  | 0.903 |  |  |  | 1.992 |  |  |  |  | 6.992
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC | x Chan | 74 Degrees |  | Added May 2004 | [11 empty] | File-74oa3159.dat
Low? Gain | 9.98 | 2960 | 0.903 | 0.4515 | 23.2 | 0.062703 | -38.7493 |  | 8.0259 |  | 74oa3159.dat |  | -35.2605 |  |  | Index=1.45
| 9.98 | 2960 | 1.992 | 0.996 | 75 | 0.2027 | -20.9973 |  | 25.9459 |  | 74oa3159.dat |  | -16.6405 |  | 1 | 58.6 | 0.15838
| 9.98 | 2960 | 3.4 | 1.7 | 183 | 0.49459 | 16.0146 |  | 63.3081 |  | 74oa3159.dat |  | 22.1811 |  | 3 | 357 | 0.96486
| 9.98 | 2960 | 4 | 2 | 247 | 0.66757 | 37.9476 |  | 85.4486 |  | 74oa3159.dat |  | 45.1865 |  | 5 | 887.5 | 2.3986
| 9.98 | 2960 | 4.92 | 2.46 | 332 | 0.8973 | 67.0773 |  | 114.8541 |  | 74oa3159.dat |  | 75.7405 |  | 7 | 1530 | 4.1351
| 9.98 | 2960 | 6.992 | 3.496 | 699.2 | 1.8897 | 192.9177 |  | 241.8854 |  | 74oa3159.dat |  | 207.7341 |  | 9.98 | 2960 | 8
0
0
OPC | x Chan | 74 Degrees | Trude's numbers |  | Checked Nov 2005 |  |  |  | 0 | [6 empty] | File-74oa3159.dat
Low? Gain | 5 | 887.5 | 0.903 | 0.4515 | 23.2 | 0.20913 | -20.1827 |  | 26.7682 |  | 74oa3159.dat |  | -15.7861 |  |  | Index=1.45
| 5 | 887.5 | 1.992 | 0.996 | 75 | 0.67606 | 39.0239 |  | 86.5352 |  | 74oa3159.dat |  | 46.3155 |  | 1 | 58.6 | 0.15838
| 5 | 887.5 | 3.4 | 1.7 | 183 | 1.6496 | 162.4664 |  | 211.1459 |  | 74oa3159.dat |  | 175.7938 |  | 3 | 357 | 0.96486
| 5 | 887.5 | 4 | 2 | 247 | 2.2265 | 235.6175 |  | 284.9893 |  | 74oa3159.dat |  | 252.5217 |  | 5 | 887.5 | 2.3986
| 5 | 887.5 | 4.92 | 2.46 | 332 | 2.9927 | 332.7713 |  | 383.0625 |  | 74oa3159.dat |  | 354.4259 |  | 7 | 1530 | 4.1351
| 5 | 887.5 | 6.992 | 3.496 | 699.2 | 6.3026 | 752.4758 |  | 806.7389 |  | 74oa3159.dat |  | 794.6522 |  | 9.98 | 2960 | 8
0
0
OPC | x Chan | 74 Degrees | High gain option | [5 empty] | 0 | [6 empty] | File-74oa3159.dat
Low? Gain | 1 | 58.6 | 0.903 | 0.4515 | 23.2 | 3.1672 | 354.9055 |  | 405.4061 |  | 74oa3159.dat |  | 377.6423 |  |  | Index=1.45
| 1 | 58.6 | 1.992 | 0.996 | 75 | 10.2389 | 1251.5935 |  | 1310.5802 |  | 74oa3159.dat |  | 1318.1747 |  | 1 | 58.6 | 0.15838
| 1 | 58.6 | 3.4 | 1.7 | 183 | 24.9829 | 3121.1362 |  | 3197.8157 |  | 74oa3159.dat |  | 3279.1304 |  | 3 | 357 | 0.96486
| 1 | 58.6 | 4 | 2 | 247 | 33.7201 | 4229.0133 |  | 4316.1775 |  | 74oa3159.dat |  | 4441.1782 |  | 5 | 887.5 | 2.3986
| 1 | 58.6 | 4.92 | 2.46 | 332 | 45.3242 | 5700.4126 |  | 5801.5017 |  | 74oa3159.dat |  | 5984.5229 |  | 7 | 1530 | 4.1351
| 1 | 58.6 | 6.992 | 3.496 | 699.2 | 95.4539 | 12056.8577 |  | 12218.1024 |  | 74oa3159.dat |  | 12651.772 |  | 9.98 | 2960 | 8

## 74 Deg Cal Points
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC | x Chan | 74 Degrees |  | Added May 2004 | [7 empty] | File-74oa3159.dat
Low? Gain | 9.98 | 2960 | 0.903 | 0.4515 | 23.2 | 0.062703 | -38.7493 | 8.0259 | 74oa3159.dat | -35.2605 |  | Index=1.45
| 9.98 | 2960 | 1.992 | 0.996 | 75 | 0.2027 | -20.9973 | 25.9459 | 74oa3159.dat | -16.6405 | 1 | 58.6 | 0.15838
| 9.98 | 2960 | 3.4 | 1.7 | 183 | 0.49459 | 16.0146 | 63.3081 | 74oa3159.dat | 22.1811 | 3 | 357 | 0.96486
| 9.98 | 2960 | 4 | 2 | 247 | 0.66757 | 37.9476 | 85.4486 | 74oa3159.dat | 45.1865 | 5 | 887.5 | 2.3986
| 9.98 | 2960 | 4.92 | 2.46 | 332 | 0.8973 | 67.0773 | 114.8541 | 74oa3159.dat | 75.7405 | 7 | 1530 | 4.1351
| 9.98 | 2960 | 6.992 | 3.496 | 699.2 | 1.8897 | 192.9177 | 241.8854 | 74oa3159.dat | 207.7341 | 9.98 | 2960 | 8
0
0
OPC | x Chan | 74 Degrees | Trude's numbers |  | Checked Nov 2005 |  |  | 0 |  |  |  | File-74oa3159.dat
Low? Gain | 5 | 887.5 | 0.903 | 0.4515 | 23.2 | 0.20913 | -20.1827 | 26.7682 | 74oa3159.dat | -15.7861 |  | Index=1.45
| 5 | 887.5 | 1.992 | 0.996 | 75 | 0.67606 | 39.0239 | 86.5352 | 74oa3159.dat | 46.3155 | 1 | 58.6 | 0.15838
| 5 | 887.5 | 3.4 | 1.7 | 183 | 1.6496 | 162.4664 | 211.1459 | 74oa3159.dat | 175.7938 | 3 | 357 | 0.96486
| 5 | 887.5 | 4 | 2 | 247 | 2.2265 | 235.6175 | 284.9893 | 74oa3159.dat | 252.5217 | 5 | 887.5 | 2.3986
| 5 | 887.5 | 4.92 | 2.46 | 332 | 2.9927 | 332.7713 | 383.0625 | 74oa3159.dat | 354.4259 | 7 | 1530 | 4.1351
| 5 | 887.5 | 6.992 | 3.496 | 699.2 | 6.3026 | 752.4758 | 806.7389 | 74oa3159.dat | 794.6522 | 9.98 | 2960 | 8
0
0
OPC | x Chan | 74 Degrees | High gain option |  |  |  |  | 0 |  |  |  | File-74oa3159.dat
Low? Gain | 1 | 58.6 | 0.903 | 0.4515 | 23.2 | 3.1672 | 354.9055 | 405.4061 | 74oa3159.dat | 377.6423 |  | Index=1.45
| 1 | 58.6 | 1.992 | 0.996 | 75 | 10.2389 | 1251.5935 | 1310.5802 | 74oa3159.dat | 1318.1747 | 1 | 58.6 | 0.15838
| 1 | 58.6 | 3.4 | 1.7 | 183 | 24.9829 | 3121.1362 | 3197.8157 | 74oa3159.dat | 3279.1304 | 3 | 357 | 0.96486
| 1 | 58.6 | 4 | 2 | 247 | 33.7201 | 4229.0133 | 4316.1775 | 74oa3159.dat | 4441.1782 | 5 | 887.5 | 2.3986
| 1 | 58.6 | 4.92 | 2.46 | 332 | 45.3242 | 5700.4126 | 5801.5017 | 74oa3159.dat | 5984.5229 | 7 | 1530 | 4.1351
| 1 | 58.6 | 6.992 | 3.496 | 699.2 | 95.4539 | 12056.8577 | 12218.1024 | 74oa3159.dat | 12651.772 | 9.98 | 2960 | 8

## History of modifications
August, 2002 | To add new poly sizes from current batch replacing calibration sizes which were finished
April, 2003 | To add calculations for new poly sizes just purchased. The lines in blue are the new sizes.
Spread sheet also updated to account for the source file for the CR calculations, either 40253159.dat or nuwy_opc.dat.
Nuwy_opc.dat is the file produced with the latest version of the Counter response calculation
Also added the discriminator level calculation in the far right columns using both 40253149 and nuwy_opc. Dat
May, 2004 | Calculations added for 74 degree unit.
May, 2006 | Adjusted the ranges of the mid and low gain boards for calibrating the new bulb instrument.

## LabTests_New_Bulb_March_2006
Last modified April 2003 (td) |  |  |  | Calibration constants from new poly in Blue | [5 empty] | Calibration constants to use in bold
Counter | Max Radii | Cntr Respons | Poly | Poly | CntrRespon | Voltage | Channel# |  | Channel# |  | CR Calc |  | Channel#
Board | Meas by | Mie Calc | Diam | Radius | Mie Calc |  | QunMiao2 = |  | Old style = |  | Source |  | QunMiao =
| Board (um) | m-1.45 |  |  | m=1.59 |  | 126.8*V-46.7 |  | 128*V |  |  |  | 133*V-43.6
[7 empty] | Current |  | Previous |  |  |  | Used only |  | Channel | CntrRespon | Disc Level
[7 empty] | Conversion |  | convesion |  |  |  | briefly |  | Radius | m=1.45 | V
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
,--------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------------
OPC_PSC | 12 Chan | 40 Degrees | [13 empty] | File-nuwy_opc.dat
High Gain | 1.08 | 200 | 0.903 | 0.4515 | 49.3 | 1.972 | 203.3496 |  | 252.416 |  | nuwy_opc.dat |  | 218.676 |  | 0.15 | 4.94 | 0.1976
[8 empty] | CR error | [6 empty] | 1.08 | 200 | 8
Measured |  |  | 0.5 | 0.25 | 31.8809 | 1.2752 | 115 | 24.1223
March, 2006 |  |  | 0.903 | 0.4515 | 49.2311 | 1.9692 | 203 | 27.8316
|  |  | 1.01 | 0.505 | 59.0891 | 2.3636 | 253 | 19.4184
Mid Gain | 1.75 | 565 | 1.992 | 0.996 | 210.4 | 2.9791 | 331.0518 |  | 381.3267 |  | nuwy_opc.dat |  | 352.6223 |  | 0.3 | 39.7 | 0.56212
Measured |  |  | 0.945 | 0.4725 | 47.7331 | 0.67587 | 39 | 33.4132
March, 2006 |  |  | 1.992 | 0.996 | 210.3712 | 2.9787 | 331 | 37.5441
|  |  | 2.92 | 1.46 | 355.7428 | 5.0371 | 592 | 61.6386 | [6 empty] | 1.75 | 565 | 8
Low Gain | 10 | 16900 | 6.992 | 3.496 | 2195 | 1.0391 | 85.052 |  | 132.9988 |  | nuwy_opc.dat |  | 94.5941 |  | 2.5 | 1140 | 0.53964
Measured |  |  | 4.92 | 2.46 | 1211.1889 | 0.57334 | 26 | 448.5885
March, 2006 |  |  | 6.992 | 3.496 | 2194.1345 | 1.0386 | 85 | 271.1852 | [6 empty] | 10 | 16900 | 8
|  |  | 9.9 | 4.95 | 3810.1636 | 1.8036 | 182 | 458.8155
Poly tests of new light bulb - Lou King, March 2006
Board
| Diam | Radius | Counter response | Error on Counter Response | Normalized peak | Actual/Measured peak | Calibration point |  | Error |  |  | Normalized - error | -50% | Peak | 50% | Pos error normalized
High | 0.5 | 0.25 |  |  | 115.2714 | 113 |  |  | 0.75664 |  |  | 0.67257 | 37 | 113 | 208 | 0.84071
| 0.903 | 0.4515 |  |  | 203 | 199 | 203 |  | 0.56533 |  |  | 0.66332 | 67 | 199 | 292 | 0.46734
| 1.01 | 0.505 |  |  | 252.9849 | 248 |  |  | 0.32863 |  |  | 0.34274 | 163 | 248 | 326 | 0.31452
Mid | 0.945 | 0.4725 |  |  | 39.056 | 40 |  |  | 0.7 |  |  | 0.425 | 23 | 40 | 79 | 0.975
| 1.992 | 0.996 |  |  | 331 | 339 | 331 |  | 0.17847 |  |  | 0.16224 | 284 | 339 | 405 | 0.19469
| 2.92 | 1.46 |  |  | 591.6991 | 606 |  |  | 0.17327 |  |  | 0.20792 | 480 | 606 | 690 | 0.13861
Low | 4.92 | 2.46 |  |  | 25.7865 | 27 |  |  | 0.37037 |  |  | 0.44444 | 15 | 27 | 35 | 0.2963
| 6.992 | 3.496 |  |  | 85 | 89 | 85 |  | 0.1236 |  |  | 0.10112 | 80 | 89 | 102 | 0.14607
| 9.9 | 4.95 |  |  | 182.4157 | 191 |  |  | 0.12042 |  |  | 0.13613 | 165 | 191 | 211 | 0.10471